## eSTAR 5.1 Cybersecurity Notes
Description | Attachment/Data to be Entered | Notes
Firmware & Cybersecurity/Interoperability
Cybersecurity When not considering mitigation measures for the cybersecurity risks, could your device(s)/ system result in either of the following: • multi-patient harm (resulting in temporary or medically reversible adverse events), or • serious adverse events or death in the event your device(s)/system is compromised? | Yes | Multi-patient harm refers to patient harm that can impact multiple patients either simultaneously or in rapid succession. When devices are capable of connecting (wired or wirelessly) to another medical or non-medical product, to a network, or to the Internet, there is the possibility that multiple devices can be compromised simultaneously. Because of this connectivity, if a device is compromised, or if a non-device function that could impact the device function is compromised, the device may introduce a safety risk to patients through security risk.
Risk Assessment Report | 1) Third-party assessed security risk management report | Attach your security risk management report detailing a separate, parallel,and interconnected security risk management process. This is different from your safety risk management process. Provide a security risk management report as recommended in Section V.A of the Premarket Cybersecurity Guidance such as that described in AAMI TIR57 (Principles for medical device security—Risk management) (refer Attachment 1), which includes either directly or by reference, the following items: • Risk analysis, mitigations, and design considerations pertaining to cybersecurity risks, • A traceability matrix of security risks to security controls, • Traceability to the verification reports for documented security controls, • A description of when and how security updates/patches will be provided, and • A description of the steps taken to assure devices will be delivered malware-free. In addition, the security risk management report should: • Summarize the risk evaluation methods and processes, • Detail the residual risk conclusion from the security risk assessment, • Detail the risk mitigation activities undertaken as part of a manufacturer’s risk management processes, and • Provide traceability between the threat model, cybersecurity risk assessment, SBOM, and testing documentation as well as other relevant cybersecurity risk management documentation.
Risk Assessment - Threat Model | 2) Third-party invoked threat model | Attach your threat model addressing all the end-to-end elements of the system. Manufacturers should provide threat modeling for the entire medical device system including all end-to-end connections into and/or out of the system as recommended in Section V.A.1 of the Premarket Cybersecurity Guidance (refer Attachment 2). End-to-end elements includes aspects that might not be part of the medical device such as hospital network, exempted medical device functionality per 21st Century Cures Act, Cloud infrastructure, 'other functions' as described in FDA guidance document "Multiple Function Device Products: Policy and Considerations," update infrastructure, etc. The Threat Model documentation should: • Identify medical device system risks and mitigations as well as inform the pre- and post-mitigation risks considered as part of the cybersecurity risk assessment, • State any assumptions about the medical device system or environment of use (e.g., hospital networks are inherently hostile, therefore manufacturers are recommended to assume that an adversary controls the network with the ability to alter, drop, and replay packets), and • Capture cybersecurity risks introduced through the supply chain, manufacturing, deployment, interoperation with other devices, maintenance/update activities, and decommission activities that might otherwise be overlooked in a traditional safety risk assessment process. A rationale for the methodology(ies) selected should be provided with the threat modeling documentation.
List the Threat Methodology (e.g. STRIDE, Attack Trees, Kill Chain, DREAD) that you used | TBD | There are numerous methodologies available to perform threat modeling. The methodology selected will likely depend on a multitude of factors including the device itself, environment of use, team performing the assessment, and strengths and weaknesses of a given methodology. Multiple threat modeling methodologies may be combined.
Does the threat model documentation include Architecture Views (Global System View, Multi-Patient Harm View, Updateability/Patchability View, and Security Use Case Views)? | TBD | Architecture Views are described in Section V.B.2. and Appendix 2 of the Premarket Cybersecurity Guidance. Architecture Views can either be included within the Threat Modeling documentation or as standalone documents.
Risk Management - Cybersecurity Risk Assessment | 3) Cybersecurity risk assessment methodology with acceptance criteria | The Cybersecurity Risk Assessment should be a separate risk assessment from the safety risk management process and capture the risks and controls identified from the threat model. Identified risks should be assessed according to the level of risk posed from the device and the system in which it operates. Acceptance criteria for cybersecurity risks should carefully consider the total product lifecycle (TPLC) of the medical device system. The methods used for scoring the risk pre- and post-mitigation and the associated acceptance criteria as well as the method for transferring security risks into the safety risk assessment process should be provided. Security risks and controls should be assessed for residual risks and address the fact that cybersecurity-related failures can occur either intentionally or unintentionally. Known vulnerabilities should be assessed as reasonably foreseeable risks. Also, the assessment of vulnerabilities identified during cybersecurity testing should consider the TPLC of the device as the exploitability of the vulnerability is likely to increase over the device lifecycle. Furthermore, vulnerabilities identified in Cybersecurity and Infrastructure Security Agency (CISA) Known Exploited Vulnerabilities Catalog (www.cisa.gov/known-exploited-vulnerabilities-catalog) should be designed out of the device, as they are already being exploited and expose the medical device system and users to the risk.
Please cite the page number(s) of the attachment directly above where you describe your methodology and your acceptance criteria | TBD | The submission documentation should clearly define the methods used for scoring the risk pre- and post-mitigation and the associated acceptance criteria as well as the method for transferring security risks into the safety risk assessment process. For additional guidance on cybersecurity risk assessments, refer to AAMI TIR57 Principles for medical device security—Risk management. Multiple different methodologies for assessing cybersecurity risk which assess likelihood through exploitability. Some examples are listed below however other acceptable options exist: • Common Vulnerability Scoring System (CVSS) • IEEE 11073-40101-2020 (modified CVSS) • NIST SP 800-30 • DREAD • Custom (developed by manufacturer)
Does the Cybersecurity Risk Assessment avoid using probabilities for the likelihood assessment and use exploitability instead? | TBD | Cybersecurity risks are difficult to predict, meaning that it is not possible to assess and quantify the likelihood of an incident occurring based on historical data or modeling (also known as a "probabilistic manner"). This non-probabilistic approach is not the fundamental approach performed in safety risk management under ISO 14971 and further underscores why safety and security risk management are distinct but connected processes. Instead, security risk assessment processes focus on exploitability, which is the feasibility or ease and technical means by which the vulnerability can be exploited by a threat. There are no scientific methods to reasonably assign probabilities (i.e., there is no 1/1,000,000 cybersecurity risk). The exploitability assessment relies on assessing factors which can make an exploit more/less likely. Therefore, a Cybersecurity Risk Assessment for a given risk/vulnerability on the medical device is based on: • Severity of Patient Harm (if the vulnerability were to be exploited) – Same as Safety Risk Management • Exploitability (non - deterministic)
Risk Management - Software Bill of Materials (SBOM) | 4) SBOM of SnapshotBIO | Refer Attachment 3
Support on Software Components | 5) SOUP and COTS Validation | Manufacturers may choose to provide these additional elements as part of the SBOM, or they may provide it separately, such as in an addendum. The FDA Guidance "Off-the-Shelf Software Use in Medical Devices" (refer Attachment 4) provides additional recommended submission content for OTS software used in the medical device.
List the supported operating system(s) and associated version(s) your device(s)/system uses. Be aware that if you list any operating systems that are no longer supported (e.g. Windows 7, Mac OS 9) or nearing end of support, this will be considered an inaccurate response. Type "N/A" if your device(s) does not use an operating system. | TBD | The FDA Guidance "Off-The-Shelf (OTS) Software Use in Medical Devices" recommends manufacturers establish appropriate mechanisms for assuring the continued performance, maintenance and support of the OTS Software should the original OTS software developer either make changes to the OTS software and/or terminate their support of the OTS software. Operating systems include, but are not limited to, commercial computer (e.g. Windows, Linux, Mac OS), mobile device (Android, iOS), and real-time operating systems.
Attachment | 6) Safety and Security assessment of cybersecurity vulnerabilities in the component software and a description of any controls | Manufacturers should identify all known vulnerabilities associated with the device and the software components, including those identified in CISA’s Known Exploited Vulnerabilities Catalog (www.cisa.gov/known-exploited-vulnerabilities-catalog) and the NIST National Vulnerability Database (nvd.nist.gov/vuln). For each known vulnerability, manufacturers should describe how the vulnerability was discovered to demonstrate whether the assessment methods were sufficiently robust. For components with known vulnerabilities, device manufacturers should provide in premarket submissions: • A safety and security risk assessment of each known vulnerability (including device and system impacts); and • Details of applicable safety and security risk controls to address the vulnerability. If risk controls include compensating controls, those compensating controls should be described in an appropriate level of detail.
Unresolved Anomalies - Assessment | 7) Assessment of any unresolved anomalies for cybersecurity impact | FDA guidance document "Content of Premarket Submissions for Device Software Functions" (refer Attachment 5) recommends that device manufacturers provide a list of software anomalies that exist in a product at the time of submission. For each of these anomalies, device manufacturers should conduct an evaluation of the anomaly’s impact on the device’s safety and effectiveness and consult the referenced guidance document to assess the associated documentation recommended for inclusion in such device’s premarket submission. Some anomalies discovered during development or testing may have security implications and may also be considered vulnerabilities. As a result, manufacturers should also assess the potential security impacts of anomalies. The assessment should also include consideration of any present Common Weakness Enumeration (CWE) categories (cwe.mitre.org).
Cybersecurity Metrics | 8) Data from monitoring cybersecurity metrics | Manufacturers should ensure they have appropriate processes and resources to identify, assess, and mitigate cybersecurity vulnerabilities as they are identified throughout the supported device lifecycle. To demonstrate the effectiveness of a manufacturer’s processes, manufacturers should provide the following measures and metrics: • Percentage of identified vulnerabilities that are updated or patched (defect density), • Duration from vulnerability identification to when it is updated or patched, and • Duration from when an update or patch is available to complete implementation in devices deployed in the field, to the extent known. Averages of the above measures should be provided if multiple vulnerabilities are identified and addressed. These averages may be provided over multiple time frames based on volume or in response to process or procedure changes to increase efficiencies of these measures over time. For new devices, these metrics can reflect a prior model of the device. For modified devices, these metrics should reflect the current version(s) in use.
Cybersecurity Controls | 9) Document detailing control measures on: A) Authentication controls: B) Authorization controls: C) Cryptography controls: D) Code, data, and execution integrity controls: E) Confidentiality controls: F) Event detection and logging controls: G) Resiliency and recovery controls: H) Firmware and software update controls: | Please see the Cybersecurity guidance Section V.B.1 and Appendix 1 for more information.
Architecture Views | 10) Architecture Views in Threat Modelling | Attach a document(s) that contains a Global System View, Multi-PatientHarm View, Updatability/Patchability View, and Security Use Case Views. If theseArchitecture Views are provided in the Threat Modeling documentation provided above, please cite where each of the Architecture Views is provided. Documentation of Security Architecture Views (Global System View, Multi-Patient Harm View, Updatability/Patchability View, and Security Use Case Views) should include both diagrams and explanatory text. These diagrams and explanatory text should contain sufficient details to permit an understanding of how the assets within the medical device system function holistically within the associated implementation details. For the security architecture views, manufacturers should follow the recommendations outlined in Appendix 2 of the cybersecurity guidance when determining the level of detail to include in premarket submissions. These security architecture views should: • Identify security-relevant medical device system elements and their interfaces, • Define security context, domains, boundaries, critical user roles, and external interfaces of the medical device system, • Align the architecture with (a) the medical device system security objectives and requirements, (b) security design characteristics in order to address the identified threats, and • Establish traceability of architecture elements to user and medical device system security requirements. Such traceability should exist throughout the cybersecurity risk management documentation. The Security Use Case Views should cover various operational states of elements in the medical device system (e.g. power on, standby, transition states) and assess clinical functionality states of the medical device system (e.g. programming, alarming, delivering therapy, send/receive data, reporting diagnostic results). If a particular view sufficiently captures the risks of another view identified above, manufacturers do not have to duplicate documentation. Similarly, if threat modeling documentation sufficiently captures the view, manufacturers do not have to duplicate documentation. Additionally, if one of the views listed above is not appropriate, manufacturers should instead provide an explanation for why the view is not included in the premarket submission.
Cybersecurity Labelling | TBD | The following are examples of information that may be included in labeling to communicate relevant security information to users: • Device instructions and product specifications related to recommended cybersecurity controls appropriate for the intended use environment (e.g. anti-malware software, use of a firewall, password requirements). • Sufficiently detailed diagrams for users that allow recommended cybersecurity controls to be implemented. • A list of network ports and other interfaces that are expected to receive and/or send data. This list should include a description of port functionality and indicate whether the ports are incoming, outgoing, or both, along with approved destination end-points. • Specific guidance to users regarding supporting infrastructure requirements so that the device can operate as intended (e.g. minimum networking requirements, supported encryption interfaces). Where appropriate, such guidance should include technical instructions to permit secure network deployment and servicing, and instructions for users on how to respond upon detection of a cybersecurity vulnerability or incident. • An SBOM as specified in Section V.A.4. or in accordance with an industry accepted format to effectively manage their assets, to understand the potential impact of identified vulnerabilities to the medical device system, and to deploy countermeasures to maintain the device’s safety and effectiveness. Manufacturers should provide or make available SBOM information to users on a continuous basis. If an online portal is used, manufacturers should ensure that users have up-to-date links that contain accurate information. The SBOM should be in a machine-readable format. • A description of systematic procedures for users to download version-identifiable manufacturer-authorized software and firmware, including a description of how users will know when software is available. • A description of how the design enables the device to respond when anomalous conditions are detected (i.e. security events). This should include notification to the user and logging of relevant information. Security event types could be configuration changes, network anomalies, login attempts, or anomalous traffic (e.g., send requests to unknown entities). • A high-level description of the device features that protect critical functionality (e.g. backup mode, disabling ports/communications). • A description of backup and restore features and procedures to restore authenticated configurations. • A description of methods for retention and recovery of device configuration by an authenticated authorized user. • A description of the secure configuration of shipped devices, instructions for user-configurable changes, and identification of user-configurable changes that could increase security risk for the medical device system. Secure configurations may include end point protections such as anti-malware, firewall/firewall rules, allow lists, deny lists, security event parameters, logging parameters, and physical security detection, and resetting of credentials, among others. • Where appropriate for the intended use environment, a description of how forensic evidence is captured, including but not limited to any log files kept for a security event. Log file descriptions should include how, where, and in what format the log file is located, stored, recycled, archived, and how it could be consumed by automated analysis software (e.g. Intrusion Detection System (IDS) or Security Information and Event Management (SIEM)). • Information, if known or anticipated, concerning device cybersecurity (including components) end of support and end of life. At the end of support, a manufacturer may no longer be able to reasonably provide security patches or software updates. If the device remains in service following the end of support, the manufacturer should have a pre-established and pre-communicated process for transferring the risks highlighting that the cybersecurity risks for end-users can be expected to increase over time. • Information on securely decommissioning devices by sanitizing the product of sensitive, confidential, and proprietary data and software. Please see the Cybersecurity guidance Section VI.A for more information regarding these elements.
Cybersecurity Management Plan | 11) Cybersecurity Management Plan | To maintain the safety and effectiveness of the device after marketing authorization is achieved, manufacturers should provide a cybersecurity management plan detailing how they will identify and communicate to users’ vulnerabilities that are identified after releasing the device.
Please cite the page number(s) of the attachment directly above where you include the information specified in the help text to the right for all the recommended elements | 11) Cybersecurity Management Plan | he Cybersecurity Management Plan should include the following elements: • Personnel responsible; • Sources, methods, and frequency for monitoring and identifying vulnerabilities (e.g., researchers, NIST national vulnerability database (NIST NVD), third party software manufacturers); • Identify and address vulnerabilities identified in CISA Known Exploited Vulnerabilities Catalog (www.cisa.gov/known-exploited-vulnerabilities-catalog); • Periodic security testing; • Timeline to develop and release patches; • Update processes; • Patching capability (i.e., rate at which update can be delivered to devices); • Description of their coordinated vulnerability disclosure process; and • Description of how the manufacturer intends to communicate forthcoming remediations, patches, and updates to customers.
Please cite the page number(s) of the attachment directly above where you include a description of and justification for the time-lines to make patches on a regular cycle and out of cycle | 11) Cybersecurity Management Plan | Section 524B(b)(1) of the FD&C Act requires that manufacturers provide a plan to monitor, identify, and address, as appropriate, in a reasonable time, postmarket cybersecurity vulnerabilities and exploits, including coordinated vulnerability disclosure and related procedures. This plan should include a description of and justification for the timelines to make patches available as required by 525B(b)(2)(A-B): • 524B(b)(2)(A) requires manufacturers to make patches for known unacceptable vulnerabilities available on a reasonably justified regular cycle • 524B(b)(2)(B) requires manufacturers to make available patches for critical vulnerabilities that could cause uncontrolled risks as soon as possible out of cycle

## eSTAR 5.1 Checklist
Description | Attachment Name and Template Location | Justification and Comments | Status of the Line Item
Device Description - Listing of Devices
Listing of devices
Device Description - General Device Characteristics
Is the device life-supporting or life-sustaining?
Are there any direct or indirect tissue contacting components
Is the device or a component an implant?
Does the device use software/firmware
Is the device, or does it contain, digital health technology
Please check the attributes that are applicable to your device
Is the device or a component packaged as sterile?
The device/system uses or is... (choose all that apply)
The environment of use of the device/system includes... (choose all that apply)
Is the device a combination product or does a Request For Designation (RFD) / pre-RFD exist for the product?
Is the device electrical (battery or wall powered)?
Does the device/system include wireless technology?
Please check the attributes that are applicable to your device. If none apply, keep all unchecked.
Device Description
Summary
Comprehensive Device Description and Principles of Operation Documentation
Device Pictures, Illustrations, Schematics, and/or Diagrams. Attach a justification if the device does not have a physical form
Description of Device Packaging
System/Kit Components and Accessories
Is the device intended to be marketed with multiple system/kit components or accessories?
Under section 513(f)(6), are you requesting risk-based classification of an accessory that is not explicitly identified in a classification regulation, or has not been included in a cleared 510(k), approved PMA, or granted De Novo request?
Kit Certification Statement Declaration
Kit Components
Kit Schematics
Guidance & Special Control Adherance
In the text box below, please address or provide information demonstrating compliance with any applicable special controls, requirements in a device specific classification regulation, or adherence to device specific guidance recommendations for the Device Description. For each specific special control, regulation, or guidance recommendation applicable to your device: Please list the special control, regulation, or recommendation and cite the attachment(s) and page number(s) where it was addressed. Please use the primary product code you will provide in the Classification section below to determine whether a device specific guidance exists for your device. Type “N/A” if no device specific guidance, regulation, or special controls exist for your device type. If you type “N/A,” and special controls, a regulation, or a device specific guidance exists that requests information covered by this review section, the time-line for review of your file may be affected
Indications for Use
Submission #
Device Name
Indications for Use description
Type of Use
Classification
Medical Specialty
Regulation
Product Code
Associated Product Code(s)
Predicates and Substantial Equivalence
Primary Predicate
Is this a Preamendments or Exempt device without a submission number?
Predicate Submission Number (e.g., K210001)
Predicate Device Trade Name
Medical Specialty
Regulation
Product Code
Substantial Equivalence Comparison - Indications for Use
Substantial Equivalence Comparison - Technology Characteristics
Comparison Table
Labelling
If a symbols glossary was used, please specifically cite the attachment and page number where it is located in the labeling (type "N/A" if not used). Be aware that if a glossary was not used, the symbols should be described in adjacent text (if applicable, see Help Text).
What is the Magnetic Resonance (MR) safety status for the device(s) in the submission?
If the device(s) contains a static/permanent magnet, pleasespecifically cite the attachment and page number where labelingincludes a statement and/or contraindication regarding the risk of magnet use within 6 inches of other magnetically susceptiblemedical devices. Type "N/A" if not applicable.
Package Labelling Attachment
Package Inserts Attachment
Other Labelling Attachment
Please specifically cite the attachment and page number where the prescription statement or "Rx only" exists in the labeling.
Please specifically cite the attachment and page number where the Indications for Use exists in the labeling
Please specifically cite the attachment and page number where the name and place of business of the manufacturer, packer, or distributor is located
Guidance & Special Control Adherance
In the text box below, please address or provide information demonstrating compliance with any applicable special controls, requirements in a device specific classification regulation, or adherence to device specific guidance recommendations for the Device Description. For each specific special control, regulation, or guidance recommendation applicable to your device: Please list the special control, regulation, or recommendation and cite the attachment(s) and page number(s) where it was addressed. Please use the primary product code you will provide in the Classification section below to determine whether a device specific guidance exists for your device. Type “N/A” if no device specific guidance, regulation, or special controls exist for your device type. If you type “N/A,” and special controls, a regulation, or a device specific guidance exists that requests information covered by this review section, the time-line for review of your file may be affected
Reprocessing, Sterility, and Shelf-Life
Are cleaning or disinfection or sterilization instructions included in the labeling?
How many different sets of reprocessing instructions are included (maximum of 4)
Please identify the device(s), and/or accessory(ies), and/or component(s) that are covered by this set of instructions.
Please specifically cite the attachment(s) and page number(s) where the cleaning/disinfection instructions are located in the labeling.
Some devices (e.g., some orthopedic implants) are provided clean and are onlysubjected to end-user sterilization. Is the device single-use and provided clean so that cleaning instructions are unnecessary and are not included in the labeling?
Is a Point-of-Use care instruction included (e.g., to prevent drying)
Please specifically cite the attachment(s) and page number(s) where the disassembly/reassembly instructions are located. Type "N/A" if not applicable. (Note: If disassembly of the device is necessary for adequate reprocessing, specific instructions for disassembly and reassembly should be provided
Please specifically cite the attachment(s) and page number(s) where the instructions to “thoroughly clean” the device (or similar text) are included in the labeling? Type "N/A" if not applicable
If the labeling recommends special accessories for cleaning, please specifically cite the attachment(s) and page number(s) or section(s) where the instructions include details on special accessories (e.g., brush size, brush materials, detergent category) if applicable. Type "N/A" if not applicable
Please specifically cite the attachment(s) and page number(s) where the instructions include how to use accessories (e.g., medical washers) or reference accessory labeling. Type "N/A" if not applicable
Please specifically cite the attachment(s) and page number(s) where the instructions include specifications for cleaning (e.g., times, temperatures, cleaning agents, dilution/concentration, rinses including duration or volume, time, and appropriate final rinse water quality). Type "N/A" if not applicable.
Please specifically cite the attachment(s) and page number(s) where the instructions include: 1) visual inspection, 2) an inspection endpoint (e.g., “no visual contamination”), and 3) steps the user should repeat if the endpoint is not met.
Please specifically cite the attachment(s) and page number(s) where the instructions include mid-process drying (i.e., drying after cleaning)recommendations. Type "N/A" if not applicable
If lubricating agents are recommended, do the reprocessing instructions recommend the use of a class of lubricating agent (e.g., water soluble lubricants)after cleaning that is compatible with the medical device, its intended use, and withany subsequent processing steps such as sterilization
Please specifically cite the attachment(s) and page number(s) where the instructions include reuse life limits (e.g., number of uses, inspection specifications or performance tests). Type "N/A" if not applicable
Which types of microbicidal process instructions are present in these instructions?
Please specifically cite the attachment(s) and page number(s) where the instructions recommend use of a commonly available disinfecting agent
Guidance & Special Control Adherance
In the text box below, please address or provide information demonstrating compliance with any applicable special controls, requirements in a device specific classification regulation, or adherence to device specific guidance recommendations for the Device Description. For each specific special control, regulation, or guidance recommendation applicable to your device: Please list the special control, regulation, or recommendation and cite the attachment(s) and page number(s) where it was addressed. Please use the primary product code you will provide in the Classification section below to determine whether a device specific guidance exists for your device. Type “N/A” if no device specific guidance, regulation, or special controls exist for your device type. If you type “N/A,” and special controls, a regulation, or a device specific guidance exists that requests information covered by this review section, the time-line for review of your file may be affected
Shelf-Life: Does your device have a shelf-life
What is the proposed shelf-life?
Provide a summary of the methods used to support the sterility and performance of the device over its proposed shelf-life. Alternatively, provide a rationale for why testing to establish shelf-life is not applicable.
Attachment
Biocompatibility
How many tissue contacting products/components/materials are there?
Identify the device(s) / accessory(ies) / component(s) that directly or indirectly contacts the tissue
Please state the exact name and any identifiable information for the particular material used
If color additives are included, please identify them here. If no color additives are included, state "N/A."
Choose intended contact of the particular material
Please provide the FDA submission number (e.g., K210001) if you are aware of a previously submitted device using the same material with similar nature of contact
Are the listed components in contact with intact skin only AND are all the component materials included in Attachment G, Section B of the FDA Biocompatibility Guidance
Is there a potential for repeat exposure?
Choose the type of tissue contact of your tissue contacting material.
Duration of Contact
Attachment
Software/Firmware & Cybersecurity/Interoperability
Documentation Level Evaluation
Attachment
Software Documentation
Cybersecurity When not considering mitigation measures for the cybersecurity risks, could your device(s)/ system result in either of the following: • multi-patient harm (resulting in temporary or medically reversible adverse events), or • serious adverse events or death in the event your device(s)/system is compromised?
Risk Assessment Report
Risk Assessment - Threat Model
List the Threat Methodology (e.g. STRIDE, Attack Trees, Kill Chain, DREAD) that you used
Does the threat model documentation include Architecture Views (Global System View, Multi-Patient Harm View, Updateability/Patchability View, and Security Use Case Views)?
Risk Management - Cybersecurity Risk Assessment
Does the Cybersecurity Risk Assessment avoid using probabilities for the likelihood assessment and use exploitability instead?
Risk Management - Software Bill of Materials (SBOM)
Support on Software Components
List the supported operating system(s) and associated version(s) your device(s)/system uses. Be aware that if you list any operating systems that are no longer supported (e.g. Windows 7, Mac OS 9) or nearing end of support, this will be considered an inaccurate response. Type "N/A" if your device(s) does not use an operating system.
Attachment
Unresolved Anomalies - Assessment
Cybersecurity Metrics
Cybersecurity Controls
Architecture Views
Cybersecurity Labelling
Cybersecurity Management Plan
Interoperability: How many Electronic Interfaces are there?
Name the electronic interface
Is the electronic interface inactive (i.e. not meant to connect, exchange, or use data with or from other medical devices, products, technologies, or systems)?
EMC, Wireless, Electrical, Mechanical, and Thermal Safety
EMC: How many devices/accessories/components were subjected to EMC testing (maximum of 4)?
EMC Test 1 - Device Specific EMC Considerations Please select the highest possible risk associated with this device(s) resulting from potential malfunction, disruption, or degradation due to electromagnetic interference (EMI) that could cause harm to the patient, user, or operator
Was testing performed according to a recognized edition of IEC 60601-1-2 or (for IVDs only) IEC 61326-2-6 and, if so, was an ASCA test summary report provided? Please attach any summaries/reports in the "EMC, Wireless & EMT Documentation" section below
Performance Testing
Was Bench Testing used in order to support this submission?
Was Animal Testing used in order to support this submission?
Was Clinical Testing used in order to support this submission?
Adminstrative document
Executive summary
Attach User Fee
Declaration of conformity
Fill and sign in estar itself
Verification
Submission
Finally submit the signed estar
Note on D91: Monisha Babu: Please verify how many contact material are present in the device